### Pueblo of Santa Ana Department of Natural Resources

Quality Assurance Project Plan (QAPP) PurpleAir Ambient Air Quality Monitoring Environmental Division

Quality Assurance Project Plan PurpleAir Ambient Air Quality Monitoring

Santa Ana Pueblo Approvals:

Date:

Date:

I

Date:

EPA Approver:

'

Date:	 _

| List of Acronyms |  |
| --- | --- |
| AQI | Air Quality Index |
| AQP | Air Quality Program |
| AQRV | Air Quality Related Value |
| DQO | Data Quality Objectives |
| ED | Environmental Division |
| EPA | Environmental Protection Agency |
| FEM | Federal Equivalent Method |
| FRM | Federal Reference Method |
| ITEP | Institute for Tribal Environmental Professionals |
| PAM | PurpleAir Monitor |
| PM10 | Particulate Matter 10 micrometers or smaller |
| PM2.5 | Particulate Matter 2.5 micrometers or smaller |
| SADNR | Santa Ana Department of Natural Resources |
| TAMS | Tribal Air Monitoring Support |
| QAM | Quality Assurance Manager |
| QAO | Quality Assurance Officer |
| QAPP | Quality Assurance Project Plan |
| QMP | Quality Management Pan |
| QREST | Quality Review and Exchange System for Tribes |

Table of Contents

  - Project Organization	6

  - Distribution List	6

  - Project Organization Chart	6

  - Roles and Responsibilities	7

  - Project Overview	7

  - Air Quality Program Purpose/Background	7

- Problem Definition	8

- Problem Description	8

  - Summary of Method	8

  - Schedule Overview	9

- Data Quality Objectives and Criteria for Measuring	10

  - Stating the Problem	10

  - Identifying the Decision	10

  - Identifying the Inputs to the Decision	10

  - Defining the Boundaries of the Project	10

  - Deciding on a Decision Rule	11

  - Optimizing the Design	11

- Data Quality Indicators	11

- Special Training/Certification	12

- Documents and Records	12

  - QAPP/SOP/Other Documents	12

  - Reports	13

- Sampling Design	13

  - Installment/Connection/Sensor Registration	13

  - Site Selection and Placement	14

  - Field Data Acquisition and Management	15

  - Sample Handling and Custody	15

  - Analytical Methods	15

- Quality Control Requirements	15

- Data Operations and Management	16

  - Equipment	16

  - Software	16

  - Equipment Testing, Inspection, and Maintenance	16

  - Equipment Calibration and Frequency	17

  - Inspection/Acceptance of Supplies and Consumables	17

  - Non-Direct Measurements	17

  - Data acquisition and management	17

- Validation and Verification Methods	18

- Community Outreach	18

References	20

# Project Organization

#### Distribution List

The following individuals will receive copies of the approved Quality Assurance Project Plan and any subsequent approved revisions:

| Name | Title | Organization | Contact Information |
| --- | --- | --- | --- |
|  | Project Officer, Region 6 | USEPA |  |
|  | QA Officer, Region 6 | USEPA | TBD |
|  | SADNR Director | SADNR |  |
|  | ED Manager, QA Manager | SADNR |  |
|  | Environmental Specialist, QA Officer | SADNR |  |

The Santa Ana Pueblo Air Quality Program is a part of the Environmental Division (ED) within the Santa Ana Department of Natural Resources. The PurpleAir monitors will be operated by two staff members within this program: the Environmental Manager and the Environmental Specialist. These staff members will oversee maintaining and updating quality assurance documents, as well as the installation, maintenance, and data retrieval and analysis from the PAM.

Data collected will be quality assured and corrected by the ED before it is formally reported to the SADNR director, governors, tribal council, community, and other relevant parties.

#### Project Organizational Chart

Environmental Specialist, QAO

Environmental Program Manager, QAM

SADNR Director

  - Roles/Responsibilities

The Quality Assurance Officer (QAO) will collect and verify data from the PAM following the requirements of this QAPP. They will oversee that the monitor is functioning properly, as well as process and analyze the data. The QAO will also upload verified data to the AQP online database and QREST database. It is the QAO’s responsibility to contact the PurpleAir manufacturer as needed to fix equipment malfunctions or failure.

The Quality Assurance Manager (QAM) will review data that has been verified. It is the responsibility of the QAM to update and submit the QAPP for review by the SADNR director and EPA project officers.

The SADNR director will oversee all program activities.

## Project Overview

#### Air Quality Program Purpose/Background

The Santa Ana Department of Natural Resources (SADNR) Air Quality Program (AQP) acquired two PurpleAir monitors (PAM) in the Fall of 2020. With these monitors, the program can begin the first semi-permanent1 and continuous monitoring of ambient air on the Pueblo. These monitors measure for PM 2.5 and contain a temperature/pressure/humidity sensor.

Though these monitors have not been approved for regulatory purposes, we hope to collect reliable and quality data. The PAM use laser particle counters that are able to detect 0.3um- to 10 um-sized

1 The PurpleAir monitors have an estimated lifespan of 2-3 years before they need to be sent in to be recalibrated or replaced, as funding allows.

particles and differ from regulatory sensors by the method of measuring particulate matter and the averaging time of the data collected. Once these monitors are installed and operating, we will keep internal records, monitor data being automatically reported to the PurpleAir website, and upload data to be stored on the Quality Review and Exchange System for Tribes (QREST), an online database.

Low-cost sensors are continually becoming more popular for personal use, research, and even to augment regulatory networks. These sensors are desirable because they can provide data for more locations and at finer time scales at a much lower cost than methods used for regulatory purposes. While the data from sensors are not designed to be used for regulatory purposes, they may be of sufficient quality to be used for qualitative purposes and potentially even for some quantitative purposes, as long as data uncertainty is properly understood and communicated.

The goal of the AQP is to protect and improve the Pueblo’s ambient and indoor air quality, and to inform the tribal community about air quality. This monitoring will allow us to accomplish the following:

- Assess the ambient air quality of the Santa Ana Pueblo airshed and determine trends in PM2.5

- Monitor and measure the particulate matter as a result of various air pollution sources such as construction or dirt roads on and off the Pueblo, as well as the increased frequency and intensity of wildfires that have become a trend due to climate change

- Use data recorded to identify the need for monitoring of other air quality parameters

- Incorporate data to inform possible future development of PM2.5 tribal Air Quality Standards for the Pueblo of Santa Ana

Temporary air monitoring has occurred on the Pueblo with the use of the TAMS and the Cherokee ITEC Mobile Air Monitoring Network loaned equipment.. The main user of resultant data from the PAM will be the Pueblo of Santa Ana.

This Quality Assurance Project Plan (QAPP) provides the framework to operate air quality monitors under the air quality program. This is a working document that will be revised on an annual basis, or before there are changes in equipment.

# Problem Definition

This project was developed in response to growing concerns of the impact of PM2.5. and its impact on health and the Pueblo’s identified Air Quality Related Values (AQRVs) such as visibility, wildlife, and native plant collection.

Particulates less than or equal to 2.5 micrometers (termed “fine” particles, in contrast to the larger “coarse” PM10 particles which includes all particles less than 10 micrometers in aerodynamic diameter) are generally emitted from activities such as industrial and residential combustion and from vehicle exhaust. Fine particles are also formed in the atmosphere from gases such as sulfur dioxide, nitrogen oxides, and volatile organic compounds. These are emitted from combustion activities and then become particles as a result of chemical transformations in the air. Health effects

from inhalation include premature death and increased hospital admissions and emergency room visits (primarily among the elderly and individuals with cardiopulmonary disease); increased respiratory symptoms and disease (among children and individuals with cardiopulmonary disease such as asthma); and changes in lung tissue and structure and in respiratory tract defense mechanisms.

There are currently no PM2.5 air monitors located on Santa Ana. By installing PurpleAir monitors on the Pueblo, local sources of PM2.5 will be able to be captured and specific to this community.

# Project Description

#### Summary of Method

The measurement goal of this ambient air quality monitoring program is to estimate the concentration, in units of micrograms per cubic meter (µg/m3), of particulate matter with mean aerodynamic effective diameter less than 2.5 microns in the ambient air characteristic of the air breathed on tribal land. The method used is an automated system that provides readings every 15 minutes, 24 hours a day, using a wifi-connected PurpleAir PA-II monitor that uploads readings to the PurpleAir website.

Data from the continuous monitor will be downloaded and reviewed on a biweekly (every two weeks) schedule or more frequently, as thoroughly described in subsequent sections of this QAPP.

The ambient air monitoring stations will be installed in two of the Pueblo villages: Rebahene and Ranchitos (see figure below). Specific information demonstrating how the locations of the monitors meets our objectives, and the rationale for the locations can be found in Section 9.

Figure 1: Map of PurpleAir monitor locations

#### Schedule Overview

After approval of this QAPP, the first measurements will be made in accordance with the requirements and limits described in Section 5. The results of QC check will be entered into our database, and calculated to determine whether QC results are within limits and identify out-of-limits results. All valid hourly PM2.5 data must be bracketed by passing QC checks (before and after, in accordance with the schedule in Sections 10 and 11). The hourly data will be downloaded at least every two weeks (biweekly) and all data will be reviewed as it is imported into the database.

At least every quarter, the PM2.5 data will undergo the final data validation process in accordance with our validation methods detailed in Section 12 . Assessments will be conducted as frequently as necessary to meet the project objectives.

Both the PM2.5 values as well as the supporting QC results will be uploaded to QREST quarterly, and reports issued to the granting agency at least every quarter.

# Data Quality Objectives and Criteria for Measuring

#### Stating the Problem

There has been concern recently that air quality conditions are being increasingly impacted on tribal land. Tribal members have voiced concern for how their air quality and AQRVs are being affected by activities both on and off the Pueblo. Some sources whose impact might be noted are from the nearby highways that run parallel to the southern and western reservation boundaries, construction activities adjacent to the Pueblo, the numerous unpaved roads across the Pueblo, the occurrence of wildfires and the subsequent smoke, and PM2.5 emissions from local permitted sources (American Gypsum Plant, Vulcan Gravel Mine, etc). Santa Ana is not currently monitoring for PM, and there are no monitors near enough to characterize the PM within the Pueblo community. If we conduct this monitoring and conditions continue to change, we will be able to track changes in air quality.

This will be especially important to inform actions that may help to improve air quality, e.g. paving roads, spreading information on better burning practices, implementing a woodstove change out program.

#### Identifying the Decision

First, the results will be used to calculate an estimated air quality index, and as relevant based on the results, to generate health advisories to the Pueblo using the AQI recommendations, especially to vulnerable population groups such as children and elders. Our goal is to generate information that can be used on an ongoing basis to provide information to the public to guide decisions on outdoor activities for vulnerable populations.

Second, we will assess the results to decide whether additional monitoring should be pursued. This may be through the purchase of additional PAM or regulatory monitor(s).

Third, the AQP will use the data to establish a baseline, so that we can track changes in air quality over time. We are very interested in how air quality changes as there are continued and active development in our community and the surrounding area.

#### Identifying the Inputs to the Decision

The type of data needed is defined by the intended use of the data, and the inputs to the decision are the data that are gathered. The data that will be gathered will be in accordance with all EPA recommendation and industry practice in terms of schedule, siting, and QC, so that the data will be adequate to compare to other monitoring locations. Data may be used by our office to compare our PM2.5 measurements with the NAAQS after it is validated.

#### Defining the Boundaries of the Project

The spatial boundary is determined by the sites where the monitors are placed, with the associated assumptions (see Section 9) about the representativeness of these locations.

#### Deciding on a Decision Rule

Consistent with our objective to provide information to the community, we will evaluate daily averages as we gather the data, and issue public health advisories based on the AQI if results indicate that any population is at risk. This would be posted to the AQP website automatically via the PurpleAir widget. The program has the future goal to set up an automatic text-alert system to inform tribal members of hazardous AQI days, with levels above 150.

#### Optimizing the Design

The design has been optimized to fit the budget and the needs of the tribe. Priority has been planned on the objectives presented and using a method that meets the requirements of measurements to be considered reliable. If further information is gathered, or if the situation changes (such as development of housing or tribal buildings in a different area) then the plan will be changed and this QAPP will be revised and reissued for review and approval.

# Data Quality Indicators

#### Precision

Precision is defined as the degree to which a set of observations or measurements of the same property, obtained under similar conditions, conform to each other. Precision is usually expressed as standard deviation, variance, or range, in either absolute or relative terms.

#### Completeness

Completeness is a measure of the amount of valid data compared to the amount that was expected to be obtained under correct, normal conditions, and can be calculated:

% Completeness = [(# of valid results)/ (# of potential samples)] x 100

Our program goal for completeness is 95% or greater. For our PAM, a valid day’s worth of data, representing 24 hours, is at least 23 hours of data.

#### Comparability

Comparability expresses the measure of confidence that one data set can be compared to another and can be combined for decision making. Comparability with other datasets of continuous PM2.5 data is assured by carefully following the manufacturer recommendations, including QC checks, and excluding data that does not meet the QC requirements. The measurements reported by the two Pueblo-owned monitors should be comparable between each other. PAM owned and operated in the immediate area (within 50 miles of the Pueblo residential area) may report measurements that are generally related with the potential to show varying values, but comparable overall patterns.

#### Accuracy

Accuracy is defined as the degree of agreement between an observed value and an accepted reference vale. Accuracy includes a combination of random error (precision) and systematic error (bias) components that are due to sampling and analytical operations. Accuracy of the sensors will be evaluated by comparing the variability within the sensor, comparing channel A to channel B laser counters (24-hr averages should be ≥±5 µg m-3 between A and B channels).

# Special Training/Certification

Workshops and courses hosted by ITEP and other similar resource agencies will be made available to project personnel. Records on personnel qualifications and training are maintained in personnel files and are accessible for review during audit activities. Adequate education and training are integral

to any monitoring program that strives for reliable and comparable data. Training is aimed at increasing the effectiveness of employees and the Santa Ana Department of Natural Resources.

Sufficient time will be provided by management to the personnel directly involved in this project to read and understand this QAPP and the reference documents. The AQP will also monitor the availability of the training courses offered by EPA’s Air Pollution Training Institute and Region 6 facilities, Northern Arizona University’s Institute for Tribal Environmental Professionals (ITEP), and private consulting firms. Such institutions conduct professional services and ensure certification of their courses offered. When circumstances warrant, staff members may be enrolled in one or more training courses offered by these institutions.

Additionally, those actively involved in this monitoring program must have completed the following Quality Assurance courses offered by EPA:

- Introduction to EPA Quality System Requirements

- Introduction to Quality Management Plans (QMPs)

- Introduction to Quality Assurance Project Plans (QAPPs), and

- Introduction to Data Quality Objectives (DQOs)

The operation of the PAM is intended for use by the general public and are considered user-friendly. Outside of the basic support from the PAM manufacturer, there is no additional required training for operating the PAM. These monitors do not require extensive upkeep or frequent calibration, outside of their recommended 2-3 year lifespan when they need to be sent in to PurpleAir for recalibration.

# Documents and Records

It is critical that management understand properly documenting the project’s activities takes time. The air monitoring program is committed to fully document all activities relating to data collection, analysis, validation, and reporting. This is a time commitment that must be factored into the responsibilities of the Environmental Program Manager/QAM and the Environmental Specialist/QAO.

#### QAPP/SOP/Other Documents

Copies of the PurpleAir user manuals, QAPP, SOP, and other procedures/protocol will be kept as electronic copies in the S: drive (Divisions>Environmental>Air Quality). Up to date hard copies will also be kept in the Environmental Division offices. Additionally, there will be network backups performed weekly.

Any calibration or maintenance records for the monitors will be kept in the Environmental Division offices. Since data is automatically recorded to the PAM website, there is no field data notebook for PAM air quality monitoring activities.

#### Reports

Reports generated from the PAM data will be prepared quarterly, with a comprehensive report prepared annually, at the end of the grant cycle.

# Sampling Design

This section describes the process of installing the PAM and the rationale for the location of their placement.

#### Installment/Connection/Sensor Registration

The following is the guidance provided by the PurpleAir website for proper installation (PurpleAir, 2021):

    - Choose a location that is convenient to reach, has access to a power outlet, and falls within range of the wifi network.

    - If possible, mount the sensor away from vents, local sources of pollution such as BBQs, and any foliage that would increase the likelihood of insects getting inside the laser counters.

    - Install the sensor in the vertical position, with the open end facing toward the ground. The housing is designed to protect the device from the elements while allowing air to flow freely past the two laser counters.

    - Be sure to mount the sensor high enough off of any surface that would allow rainwater to splash up into the underside of the sensor.

    - The power supply should be mounted so that it will not be submerged in water or covered by snow.

    - Use either cable ties or a screw to mount the sensor and power supply, and fashion a "drip loop" to prevent water from running down the wires and into the electronics.

    - PurpleAir outdoor sensors can withstand direct sun without being damaged. Mounting the sensor in a shady spot will produce temperature readings that are not affected by direct sunlight.

Please note that temperature readings are already elevated by as much as 10 degrees Fahrenheit due to heat generated by the wifi module inside the sensor.

    - Connect the power supply to a power outlet and tuck the wires away. Connecting the sensor to wifi:

      - It’s best to configure your sensor to wifi with the router, computer/phone/tablet, and sensor in the same room. After plugging in your device and confirming it is receiving power (look for the dim red glow inside the sensor housing, or for indoor sensors, make sure the device itself is glowing), on a wifi-enabled device, open the list of available wifi networks.

      - Connect to the sensor's network, PurpleAir-****. The **** is a 2-4 character code determined by your sensor. (It may take up to up to 10 minutes after the sensor is plugged in for this network to appear.)

      - Depending on your operating system or device, you may get a pop-up window or a message to sign in to the network. If you receive a message to sign in to the network, press it to make the pop-up window appear. The pop-up window will list all available wifi networks.

Note: If the pop-up window does not appear, after making sure you've selected the "PurpleAir-****" network and waiting a bit, if you still don't get the pop-up, try opening a web browser and loading a webpage. If you are indeed connected to the PurpleAir network, this will load the sensor interface with the wifi settings for the sensor. Another option is to open a webpage and in the address bar, enter the default IP address for the sensor: http://192.168.4.1/config.

      - Choose the wifi network you would like the sensor to connect to, enter the password for that network, and click Save. A hidden wifi network can be entered by selecting the option at the bottom of the list.

      - The wifi status bar will turn green once your PurpleAir sensor is connected to wifi and transmitting data. Once your device is configured to your local wifi network, it will no longer appear in the list of available networks - the fact that it has disappeared means that the sensor is successfully configured to wifi.

Registering the sensor on the PurpleAir Map:

- Go to www.purpleair.com/register.

- Enter the sensor's Device-ID exactly as printed on the sensor's label.

- Enter the associated email address, which is usually the email that was used to purchase the sensor. If you do not know the correct email address, please contact us for help or to associate a new email address. PurpleAir will need the Device-ID in order to help you.

- Complete the rest of the registration form. For additional help, you can view a full explanation of registration fields or read our registration tips below.

- Click register and look for a green "successful registration" notice. A registration confirmation email is sent to the Associated Email and Owner's Email. This email includes a link to view your sensor on the PurpleAir map.

#### Site Selection/Placement

Placement of the two PurpleAir monitors was determined based on proximity to the inhabited tribal community, access to wifi and power, and consideration of wind direction patterns from wind roses generated in previous studies or from other external, reliable sources. The wind data that is being collected at the Pueblo’s 22 weather stations may also eventually be used to inform future placement when/if the program purchases additional monitors.

Site security was also considered as these will be placed within the tribal neighborhood/community and could be subject to damage or tempering. Both selected sites will be within fenced areas, not readily accessible by the public.

#### Field Data Acquisition and Management

There is no need for field data acquisition as the PAM operate passively. The monitors are connected through wifi to the PurpleAir website, with readings being recorded and uploaded every 15 minutes. Data will be reviewed weekly to ensure each sensor’s channel A and B are reading similarly.

Data that is automatically uploaded to the PAM website will be downloaded and stored in the AQP database.

#### Sample Handling and Custody

The PAM monitors upload readings automatically to the PAM website, where data can be viewed and downloaded. As this upload is automatic and a softcopy, a chain of custody record is not required. Data will be uploaded to the AQS and also stored on the SADNR server, with periodic updates and verification as needed.

#### Analytical Methods

The PAM perform in situ measurements and do not produce any physical samples that require analysis. The maintenance and QC requirements are described in section 10. Calibration is described in section 11.

## Quality Control Requirements

The PAM readings uploaded to the PAM website will be checked biweekly (every two weeks) to ensure properly functioning sensors. Sensors on the PAM do not have a flow requirement and do not require temperature/pressure verifications.

The graphs of the channel A and B will be reviewed weekly to identify potential sensor malfunction or damage. Agreement between A and B channels provides confidence in measurements.

Temperature and relative humidity readings will not be used as an indicator of correlation or accuracy improvement needs or as an indicator of degradation.

The US-wide correction factor developed by the USEPA will be used to minimize error from raw data. The development of the correction factors used data aggregated to 24-hour averages. Points were removed if these 24-hr averaged A & B PM2.5 channels differed by ≥± 5µg m-3 and ≥± 62% (95% confidence interval on % error [2* standard deviation (% error)]).

The resulting equation that will be applied to raw PAM data to calculate corrected PM2.5 measurements is as follows (USEPA, 2020):

### PM2.5 corrected = 0.524* [PA_cf1(avgAB)]-0.0852* RH +5.72

PM2.5 = µg m-3

RH = Relative Humidity (%)

PA_cf1(avgAB) = PurpleAir higher correction factor data averaged from the A and B channels

## Data Operations and Management

#### Equipment

The PurpleAir monitors come with mounting screws and zip-ties. The two PurpleAir PA-II monitors are equipped to measure particulate matters 2.5 micrometers or less (PM 2.5 ). They also contain a combined temperature/humidity/pressure sensor.

#### Further detail of the PAM technology can be found on the PAM website: Laser Particle Counter:

PurpleAir uses PMS5003 and PMS1003 laser particle counters. These sensors count suspended particles in sizes of 0.3, 0.5, 1.0, 2.5, 5.0, and 10um. These particle counts are processed by the sensor using a complex algorithm to calculate the PM1.0, PM2.5, and PM10 mass concentration in ug/m3. PMS5003 and PMS1003 sensors come factory calibrated.

#### ESP8266 and Arduino:

PurpleAir sensors use an ESP8266 chip to talk to the particle counter and provide all functionality, including connecting to a WiFi network and uploading data to the cloud. This ESP8266 chip runs code developed using Arduino. PurpleAir firmware has remote update features, meaning we can modify the software and the Arduino air quality sensor device will download the new version and update itself. Each Arduino air quality sensor device checks for updates from time to time.

#### BME280 Sensor:

Each PurpleAir sensor includes a BME280 pressure, temperature, and humidity sensor for which graphs of these values are also provided on the PAM website.

As the BME280 sensor are just for interest, their values do not provide ideal sensing and these measurements will not be recorded.

#### Power Supply:

PurpleAir sensors are powered by a 5v USB power source.

#### Software

All readings from the PurpleAir monitors are linked to the PurpleAir website, that does not require a software download. There is also no software required when data is being entered into QREST.

#### Equipment Testing, Inspection, and Maintenance

Equipment will be inspected quarterly for dirt, debris, vegetation, trash, or other objects that may obstruct or impact the monitors readings. If malfunction becomes apparent through the data, the equipment will also be inspected for damage or blockage.

#### Equipment Calibration and Frequency

Regular calibration is not required for PAM operations and therefore will not be performed by AQP staff. The typical lifespan of the PAM is 2-3 years, at which time the monitor may be sent to the manufacturer for recalibration or a replacement monitor may be purchased.

#### Inspection/Acceptance of Supplies and Consumables

The PAM do not require any additional supplies or consumables and therefore does not require inspection of supplies or consumables.

#### Non-Direct Measurements

Meteorological Data

Meteorological data, such as temperature and humidity, measured by the PAM will not be collected from these monitors, as they are only provided for interest, rather than qualitative value.

Meteorological data will instead be obtained from the 22 weather stations located across the Pueblo that are operated by SADNR. National Weather Service, National Climate Data Center, and other regional climate centers may be used to provide information required when developing monitoring sites, computing corrections needed to convert form standard conditions to actual conditions, and to support analysis and monitoring efforts. These data are not reported to QREST and are identified if they are used in monitoring or reporting.

Geographic Location

Another type of data that will commonly be used is geographic information. For identifying the location of sampling sites conventional longitude and latitude coordinates, or UTMs are reported.

Historical Monitoring Information

Historical monitoring data and summary information derived from previous data may be used in conjunction with current monitoring results to calculate and report trends in pollutant concentrations. In calculating historical trends, it is important to verify historical data are fully comparable to current monitoring data. If different methodologies were used to gather the historical data, the biases and other inaccuracies must be described in trends reports based on that data.

Monitor Operation and Manufacturers’ Literature

Important information is found in the manufacturers’ literature and operating manuals. All supporting documentation will be available electronically on the system network in a resources folder. This includes this QAPP, all attachments, and instrument manuals.

#### Data acquisition and management

The process of downloading data from the PAM website is known as data acquisition, whereas the organization of the data is known as data management. Within both of these areas, quality assurance activities and data reviews must be carried out to verify the adequate quality of the data. A governing principle is that sufficient thoroughness and time taken in data management would allow someone else (e.g. an auditor) to reconstruct or final values reported (to the public, EPA quarterly reports, AQS) and traced back to the original raw data to convey evidence and rationale for the data’s validity/invalidity. This information must be available months or years after the data were gather, which requires a rigorous and consistently followed system and documentation.

There is no active collection of data required as measurements are automatically recorded on the PAM website after the monitor is registered on the PurpleAir website.

Data will be downloaded from the PurpleAir website following the instructions that are included as an attachment to this document.

# Validation and Verification Methods

This section describes how the Air Quality Program verifies and validates data collection operations. Verification is conducted as an ongoing process through careful logging of field/database operations, and evaluating both parameter and QC data as it is available (i.e., not waiting until data review is scheduled but reviewing data as it is available). Validation consists of “stepping back” from the process and evaluating whether the data we are gathering are useful for our purpose.

Data that is fed to the PurpleAir website via wifi will be verified before being uploaded to QREST. Any outlying data will be investigated for cause.

The readings from channel A and B will be compared to verify that the PAM is not misfunctioning. The readings between the two Pueblo owned PAM sensors will be compared, as well as the readings from other PAM in the area for validation

The data verification and validation is conducted by the Environmental Specialist/QAO and is reviewed by the Environmental Program Manager/QAM. Any identification of issues that require corrective action are logged in the form of memos and corrective action is conducted and documented. Such communications are logged as part of the data review process in the database logbook and validation report.

All information relevant to support the determination of validity of the data is included in the final data validation report and file package.

## Community Outreach

The data collected from the PurpleAir monitors will be linked to the SADNR Air Quality Program website so that it is easily accessible by pueblo community members (via a PurpleAir widget). This widget also communicates the current AQI for PM based on the PAM measurements. These monitors and data may also be used as a teaching tool for after school and summer youth programs. Data can be viewed or downloaded directly from the PurpleAir website by the public at any time. It is a long-term goal for the AQP to utilize this data to send text-alerts to the community when the AQI for PM reaches unhealthy levels (above 150).

A summarized report will also be included in the community newsletter and Facebook postings to keep tribal members informed of the data and assessments that result from the PAM study.

#### References:

City of Albuquerque. 2021. Air Quality Monitoring. https://www.cabq.gov/airquality/air-quality- monitoring

Montana Department of Environmental Quality. 2017. Montana Ambient Air Monitoring Program Quality Assurance Project Plan. http://deq.mt.gov/Portals/112/Air/AirMonitoring/Docu ments /2017_MT_QAPP_VOL_I _DRFT_WATERMARK.pdf

PurpleAir. 2021. Our Technology. https://www2.purpleair.com/pages/technology

PurpleAir. 2021. Startup Guide for PurpleAir Sensors. https://www2.purpleair.com/pages/install Tribal Air Monitoring Support Center. (n.d.). Technical Guidance for the Development of Tribal Air

Monitoring Programs. https://datatools.tamscenter.com/hosted_files/QAQC/QC_QAPP_ BAM_PM25_CTCR.pdf

USEPA. 2020. Air Sensors PurpleAir, AirNow Fire and Smoke Map, and their use Internationally.

[PowerPoint]. https://cfpub.epa.gov/si/si_public_record_Report.cfm?Lab=CEMM&dirEn tryId=350379

USEPA. 2021. Technical Approaches for the Sensor Data on the AirNow Fire and Smoke Map.

https://www.epa.gov/air-sensor-toolbox/technical-approaches-sensor-data-airnow-fire-and- smoke-map

Wisconsin Department of Natural Resources/Air Monitoring. 2020. Quality Assurance Project Plan for PurpleAir Sensor Study QAPP 111.0. https://dnr.wisconsin.gov/sites/default/files/topi c/AirQuality/PurpleAirSensorStudyQAPP.pdf

Attachment 1:

### PurpleAir Monitor Standard Operating Procedure

Downloading Data from the PurpleAir website:

- Go to the PurpleAir website at https://www2.purpleair.com/

- Click on “Map” at the top of the main menu

- Zoom into the Pueblo monitors

- In the “Map Data Layer” box at the lower left of the map, make sure “US EPA PM2.5 AQI” is selected, as well as “US EPA” under ‘Conversion’

- Click on the “Download data” icon at the bottom right of the screen; this will take you to the “Sensor data download tool” webpage

- Select the start and end dates, and select “60 minute” average, and check the A and B boxes next to the Pueblo monitor that data is being downloaded from

- Click on “download selected” to download both primary and secondary sets of data. This will be automatically downloaded as an excel sheet.

- Save excel data into the “PurpleAir Monitor Data” folder found on the the S: drive (Divisions>Environmental>Air Quality>PurpleAir Monitor Data) with SiteID_DateRange (MM.DD-DD.YYYY) for example: Ranchitos_10.1-15.2021

  - Data that has been downloaded should be “cleaned up” following the steps below.

After data is downloaded:

Data downloaded will be in UTC and will need to be converted to MST

- Select columns B to I by clicking and holding “B” to “I”

- Shift data over two columns, so the contents of the B column now start in column D

- Select date/time data that is still in column A

- Click on data tab, “Text to Columns”

- Select “Fixed width” and click Next

- Delete the break line between the date and time by double clicking on the line. Leave the break line between the time and UTC

- Click Next. Select “General”. Click Finish

- Select cell C2 and type in the equation “=A2-(6/24)” then enter to convert from UTC to MST

- Select the cell and hover mouse on the bottom right corner until a solid “+” appears. Click and drag down to autofill remaining date/time cells.

- Delete column with “UTC”, label columns of UTC date/time and MST date/time